ISONA Toolbox User’s Guide

ver 3.1

3.3.2006

# What is ISONA Toolbox?

ISONA (Image Sensor Noise Analysis) Toolbox is developed for analysing image noise from digital camera RAW data. In addition to Nokia mobile phone camera data, it recognizes several commercial digital camera RAW formats (thanks to Mr. Dave Coffin, http://www.cybercom.net/~dcoffin/    (referred 31.1.2006)).

Toolbox functions read in RAW data and analyse each color component (variable number of colors in different cameras) separately. The OECF and noise analysis methods are based on standards ISO 14524 and ISO 15739 and on Photon Transfer Curve method presented by James R. Janesick.

# Some guidelines for taking images

Generally the image noise depends on signal levels (exposure levels) and exposure time. It is recommended that the pictures are taken with a constant exposure time and the lighting is adjusted from complete darkness to the required exposure level. Varying both exposure time and exposure level increases uncertainty of the noise analysis. Another important parameter is the number of frames taken per each exposure level. In order to properly separate temporal noise from fixed pattern noise, a relatively large number of frames (e.g. 40) per exposure level are required. The more frames, the better statistical reliability. In ISO 15739 a noise analysis method using only eighth frames is presented, so it is possible but not recommendable to use small numbers of frames.

For noise analysis it is also important to use the lowest possible ISO setting (smallest amplification) of the camera.

Because the analysis routines read the data files in given directory in time order, the naming of the files is not critical but the exposure levels should preferably be monotonically increasing (or decreasing).

# Using toolbox without GUI

## Reading in data

The analysis functions generally read in the image data automatically when given the directory name where the files are located. When the user wants to examine the images separately, the data from other than Nokia cameras is read into MATLAB using the function readraw. For example, command

>> I = readraw('C:\Data\Canon_D30\CRW_1144.CRW','O');

would read in data from file C:\Data\Canon_D30\CRW_1144.CRW and  the resulting output matrix I is [image_height by image_width by number_of_colors] uint16 matrix with color components Red, Green_r, Blue and Green_b as layers.

Function readraw uses executable file dcraw_any based on Dave Coffin’s code. The original code is modified to output RAW binary data without any further processing. This EXE file has to be in the current working directory in order for MATLAB to be able to use it.

## Browsing single images

Color components (with or without trend) of single images can be browsed with function single_images. Example:

>> dname = 'D:\Data\Canon_D30\Constant_time\';

>> ext = 'crw'; index = 61:65; trm = 1;

>> out = single_images(dname,ext,index,trm);

This produces 2*number_of_colors figures per image.  Examples of figures are on next page. If the trendless images (local average removed) are not needed, the input argument trm is set to zero and the trendless images are not calculated. Then the command above would produce number_of_colors figures per image.

The output results are 3D matrices where color components from each plotted image are the matrix layers.

## Calculating Optoelectronic Conversion Functions (OECF)

OECF analysis method is defined by ISO 14524. The analysis function oecf assumes flat field target instead of OECF test chart. Therefore the images should be taken on exposure levels from complete darkness to resulting in maximum camera output, doubling at each step so that on the logarithmic scale the measurement points will be linearly separated. The same data is then appropriate for noise analysis on different exposure levels. Example:

>> dname = 'D:\Data\Canon_D30\Constant_time\';

>> ext = 'crw'; Ngroup = 10;

>> exposure = (1/30)*[0.001 20 32 60 120 230 490 1000 2100 4100];

>> [val,out] = oecf(dname,ext,exposure,Ngroup);

defines image file directory, D:\Data\Canon_D30\Constant_time\,  file name extension, 'crw', number of frames per exposure level, 20, and exposure levels as [lux*sec] (constant exposure time 1/60 sec, dark exposure marked as 0.001 lux). These are given as input arguments to function oecf which calculates the optoelectronic conversion function (average signal at each exposure level) for each color component and returns the logarithmic exposure scale val and the average color component values in cell array out. The results are plotted with command

>> plot_oecf(ext,val,out)

and look something like this:

In this case, the maximum exposure level (≈ 137 lux*sec) is not enough for reaching the camera output maximum (212-1 = 4095). It can also be seen that the blue sensors have smaller sensitivity that red and green ones.

By default, oecf uses for analysis 64 by 64 pixel area from the centre of the color component frame. If some other pixel area is preferred, it can be defined as an additional input argument pixarea ([min_row max_row min_col max_col]):

>> pixarea = [100 200 100 200];

>> [val,out] = oecf(dname,ext,exposure,Ngroup,pixarea);

## Calculating Photon Transfer Curve (PTC)

Photon transfer curve is a log-log plot of signal noise as function of signal level. Photon transfer method is used to evaluate numerous image sensor parameters in absolute terms. Calculating it based on RAW data from digital cameras does not, however, seem to give very reliable results (for example, full well is generally not reached due to limited ADC range) but it does give some idea of the camera noise behaviour.

>> dname = 'D:\Data\Canon_D30\Constant_time\';

>> ext = 'crw'; Ngroup = 20;

>> exposure = (1/30)*[0.001 20 32 60 120 230 490 1000 2100 4100];

>> [S,std_S,std_FPN,K,RN,FWC,DR,offset] = ptc(dname,ext,exposure,Ngroup);

reads in the data and calculates average signals level, temporal (shot) noise and fixed pattern noise at each exposure level, camera gain constant K (e/DN), read noise level as electrons based on resulting K value, full well capacity based on signal level where noise reaches its maximum (which is generally not the correct FWC when RAW data is concerned) and dynamic range DR defined as maximum signal level per read noise level. Output parameter offset is dark signal offset level, minimum average signal value found. Depending on the camera settings used for dark exposure frames, it may not be the actual electrical offset level of the camera.

By default, ptc also uses for analysis 64 by 64 pixel area from the centre of the color component frame. If some other area is preferred, it can be defined as an additional input argument pixarea ([min_row max_row min_col max_col]):

>> pixarea = [201 300 161 260];

>> [S,std_S,std_FPN,K,RN,FWC,DR,offset] = ptc(dname,ext,exposure,Ngroup,pixarea);

The resulting camera gain constant is given as

>> K

K =

10.4511    1.1448    3.0000

9.9556    1.7483    3.0000

9.9116    1.7774    3.0000

9.3415    1.9221    2.0000

where the first column is average K value for Red, Green1, Green2 and Blue component correspondingly, the second column is its STD and the third column is the number of points from which the K is estimated.

The read noise is given as rms(e) for each color component,

>> RN

RN =

18.6154

17.6986

17.8058

15.7338

The resulting noise curves for each color component are plotted with command

>> plot_ptc(ext,S,std_S,std_FPN)

and they look something like this:

## Fixed pattern noise

Vertical fixed pattern noise (VFPN) is a measure of the vertical strips and banding that can be seen in an image especially at low light levels. VFPN level is mathematically described as the RMS of the column-to-column difference of the image (SMIA characterization).

Correspondingly, horizontal fixed pattern noise (HFPN) is a measure of the horizontal strips and banding that can be seen in an image especially at low light levels. HFPN level is mathematically described as the RMS of the row-to-row difference of the image.

These parameters can be calculated from average of N frames by using the function fpn. As an example we use the previous set of Canon files:

>> ave_out = ave_images(dname,ext,Ngroup);

>> [VFPN,HFPN] = fpn(4095-128,ave_out)

VFPN =

0.0013    0.0012    0.0013    0.0012

0.0016    0.0015    0.0015    0.0013

0.0034    0.0036    0.0034    0.0023

0.0128    0.0126    0.0128    0.0082

0.0438    0.0441    0.0472    0.0242

HFPN =

0.0022    0.0024    0.0022    0.0024

0.0024    0.0023    0.0024    0.0022

0.0044    0.0046    0.0045    0.0029

0.0176    0.0185    0.0183    0.0093

0.0669    0.0702    0.0695    0.0359

## Noise analysis by ISO 15739

Image noise analysis procedure defined in ISO 15739 is carried out with function noise_analysis. Input parameters are the data file directory name, file name extension, exposure levels ([lux*sec]), electrical offset of the camera and number of frames per exposure level. The analysed area is 64 by 64 pixels at the centre of the color component frame. Function calculates and returns for each exposure level the following parameters: average color component frames for that area (R_ave, G1_ave, G2_ave, B_ave), STDs of the averaged color components (s_ave), temporal noise STDs (s_temp), fixed pattern noise STDs (s_fp), pixel-by-pixel STDs of each color component (std_R, std_G1, std_G2, std_B), mean values of average color components (mC), signal-to-temporal noise ratio (SNR_temp), photoresponse non-uniformity (PRNU) at 50% signal level and the exposure level where this is reached (PRNU_lev). var_diff is the mean of all squared standard deviations of the difference images calculated by differencing each image with the average image. Noise levels are given as percentage of the maximum signal level.

For example, using the same data as before:

>> dname = 'D:\Data\Canon_D30\Constant_time\';

>> ext = 'crw'; Ngroup = 10;

>> exposure = (1/30)*[0.001 20 32 60 120 230 490 1000 2100 4100];

>> offset = 127.7232; smax = 4095;

>> [s_ave,var_diff,s_temp,s_fp,mC,SNR_temp,PRNU,PRNU_lev,out] = noise_analysis(dname,ext,exposure,offset,smax,Ngroup);

The results can be visualized with function plot_noise_result:

>> plot_noise_result(exposure,s_ave,s_temp,s_fp,SNR_temp,PRNU,…

PRNU_lev,out);

and the results look like this:

>> PRNU

PRNU =

1.2734    1.3082    1.2953    1.4781

>> PRNU_lev

PRNU_lev =

2.1357

# Using GUI

ISONA Toolbox has a graphical user interface for MATLAB 7. The GUI is started from command prompt with command

>> start_ISONA

This command initializes analysis parameter structure ISONA_params and starts ISONA_ gui. GUI looks like this:

Step 1: Select the camera type/file format.

Step 2: For multiple image analysis (OECF, PTC, ISO noise analysis, average image display, VFPN/HFPN or if you want to view several images individually), select the directory containing the images you want to analyze: either click the pushbutton “Select directory” starting the uigetdir function, or type the directory name into the space reserved for it

OR

For analyzing single images (green line imbalance, defective pixel analysis, basic analysis for single image, viewing single image or VFPN/HFPN estimate from a single image) select the file to be analyzed: click the pushbutton “Select file” starting the uigetfile function.

Step 3: Set the analysis parameters: number of frames taken per exposure level, pixel area used for analysis (default is 64 by 64 pixel area in the middle of the image), exposure times (in the example, constant exposure time of 1/60 seconds was used) and lighting levels for each exposure level as lux or cd/m^2 (these should be given in the same order than the images were taken). Notice that the decimal separator is . (dot), not , (comma). Level 0.001 means dark frame; the small offset from zero is added because of logarithmic scales used in e.g. OECF analysis. These parameters are used in OECF, PTC and ISO noise analysis.

“Trend removal” option is used when trends (local mean values) of individual images or average images are wanted. Then this option is set to “Yes”. “Exposure levels to be averaged” is used when average images are calculated. The assumption is that images of all exposure levels are in the same directory, and with Exposure level parameter you can determine from which of those levels you want to see the average images.

For defective pixel analysis by method described in SMIA characterization, the default threshold value for a defective pixel is abs(local_pixel_value – average_pixel_value) >0.15 ∙ FSD. The threshold can be set as percents of FSD in field “Defective pixel threshold value”.

Step 4: Available operations are grouped on the right. “View single images” uses either uigetfile to get the names of the images you want to observe, or picks up the name for selected single file rom the file name field. It then plots color components of each image in separate windows; i.e. four images per file. If trend removal option is “Yes”, this operation also plots images of the trendless components.

“View average images” calculates and displays averages per exposure level for each color component from the files in given directory and from given exposure levels. The averages are displayed in one figure per color component.

“Calculate OECF” calculates and displays optoelectronic conversion function from the files in given directory.

“Calculate PTC” calculates and displays photon transfer curve from the files in given directory. It also displays the parameters calculated from the PTC curve.

“ISO noise analysis” carries out the noise component analysis based on ISO 15739 standard and displays the results.

“VFPN/HFPN” analyses vertical and horizontal fixed pattern noise from either color component averages calculated from several images (as is the definition in SMIA Characterization) or from a single image, if single file is selected.

“Defective pixel analysis” uses SMIA 1.0 characterization 2.2.3 and 2.2.4 to find (possibly) defective pixels (their locations and values for each color component) from the image. It views the image color components and their histograms and outputs a structure containing locations and values for the found defective pixels in each color component.

“Basic analysis for single image” calculates the mean and STD values for each color component of the selected image. It also calculates and plots the row and column averages of each color component, showing possible trends in signal levels as a function of image column/row.